## Slide 1
EGU 2016 Vienna April 22, 2016
Infrared laser spectroscopic
gas sensing
Markus W. Sigrist
ETH Zürich, Switzerland
www.lss.ethz.ch
sigristm@phys.ethz.ch
ETH Zurich
Laser Spectroscopy and Sensing

## Slide 2
ETH Zurich
Laser Spectroscopy and Sensing
Outline
Introduction: spectroscopy, laser sources
Mid-IR VECSEL: simultaneous C1-C4 alkanes detection
QCL & QEPAS: short-lived species (HONO)
DFG & multipass cell: surgical smoke analysis
Conclusions and outlook

## Slide 3
ETH Zurich
Laser Spectroscopy and Sensing
Molecular absorption in te mid-IR
From

## Slide 4
ETH Zurich
Laser Spectroscopy and Sensing
(Broadly) tunable narrow-band mid-IR lasers
Semiconductor Lasers

## Slide 5
ETH Zurich
Laser Spectroscopy and Sensing
CW power, wavelength &
tuning ranges of EC-QCLs
Wavelength & Tuning range
of specific ICL, power 3 mW
Notes: 5

## Slide 6
Sensing of alkanes (methane, ethane, propane, butane)
ETH Zurich
Laser Spectroscopy and Sensing
Natural gas / biogas composition analysis
Energy content measurement
Fuel blending and control
Optimization of power generation
(fuel cells, gas turbines)
Monitoring of hydrocarbon leaks
at pipelines and refineries

## Slide 7
Diode-pumped lead salt VECSELs
(Vertical Extended Cavity Surface Emitting Lasers)
ETH Zurich
Laser Spectroscopy and Sensing
l ≈ 3.4 mm
Tuning range: > 150 cm-1
Pulse power: 10 mWp
Duty cycle: 0.5%

## Slide 8
ETH Zurich
Laser Spectroscopy and Sensing
Experimental setup with VECSEL, sample and reference cell

## Slide 9
ETH Zurich
Laser Spectroscopy and Sensing
Measured reference spectra of C1 – C4 alkanes
Individual gases buffered in N2
at atm. pressure
averaged over
100 scans
Total acquisition
Time: 200 s
Tuning by PZT voltage
Methane
Ethane
Propane
Butane

## Slide 10
ETH Zurich
Laser Spectroscopy and Sensing
Spectrum of mixture of C1-C3 alkanes and H2O vapor
23 ppm methane
24 ppm ethane
20 ppm propane
1.7 %vv H2O vapor
Buffer gas: N2
atm. pressure
Measurement
Fit (Hitran)
Measurement time
10 x 2 sec
Detection limit:
0.6 ppm
Absorption
J. Rey et al.: Appl. Phys. B 117, 935-030 (2014)

## Slide 11
Nitrous acid (HONO)
ETH Zurich
Laser Spectroscopy and Sensing
Important source of OH radicals in earth atmosphere
Key role in atmospheric oxidation capacity which affects regional air quality and global climate change
HONO sources and sinks not well understood due to challenging measurement:
  Atmospheric concentration only few ppb, >10 ppb indoor
- Atmospheric lifetime: 10 - 20 minutes
Sensitive and fast measurement technique is essential

## Slide 12
ETH Zurich
Laser Spectroscopy and Sensing
Short-lived species detection (HONO) with EC-QCL-QEPAS
EC-QCL (ca.1255 cm-1)
50 mW
QEPAS with mR
Compact 40 mm3 cell
Calibration with 110 m
cell and DFB QCL
Air sampling resid. time:
10 ms vs 7 min.
H. Yi, W. Chen et al.:
Appl. Phys. Lett. 106,
101109 (2015)

## Slide 13
ETH Zurich
Laser Spectroscopy and Sensing
Allan variance plot
Detection Limit 66 ppb
(1 sec) , 7 ppb (150 sec)
Good linearity of 2f QEPAS
signals with titrated HONO
concentrations
With 1W QCL: Extrapol.det. limit
3.3 ppb (1s), 330 ppt (150 s)
Poster X2.446

## Slide 14
Analysis of surgical smoke: in vivo studies
Smoke produced during minimal-invasive surgery with electro-knives or lasers. Smoke samples are taken at the hospital, collected in Tedlar bags, followed by laser and FTIR spectroscopic analysis in our lab
Collaboration with
University Hospital Zürich
(Dr. Dieter Hahnloser)
Photo of electro-knife
ETH Zurich
Laser Spectroscopy and Sensing
M. Gianella et al.: Appl Phys. B 109, 485 (2012)
M. Gianella et al.: Innov. Surgery (2013)

## Slide 15
ETH Zurich
Laser Spectroscopy and Sensing
DFG spectrometer
Broadly tunable, Mode-hop free, Room temperature
1 ECDL, 1520-1600nm
5 mW CW
2 Wavemeter for ECDL
9 Nd:YAG, 1064.5 nm,
5 kHz, 6 ns, 300 mW av
14 PPLN, 5 cm, 8 periods
23 Heatable multipass cell
up to 35 m
21/25 Detectors (VIGO)
Idler: 150 mW av.
2817 - 2920 cm-1 (29.5 mm)
2900 - 3144 cm-1 (29.9 mm)
Step size 0.002 cm-1
Minimum detectable absorption coefficient: 5×10-7 cm-1 Hz-1/2
(few ppm for many compounds of interest)

## Slide 16
ETH Zurich
Laser Spectroscopy and Sensing
Spectral analysis of surgical smoke
Principal Component Analysis (PCA)
with improved Mix-Match Algorithm
and PNNL database (with 300 gaseous compounds, but no sevoflurane)
M. Gianella et al.: Appl. Spectr. 63, 338 (2009)

## Slide 17
Chemical composition of in vivo samples
ETH Zurich
Laser Spectroscopy and Sensing
From: M. Gianella et al.: Surgical Innovation 2013 (online : June 26, 2013)
Notes: 17

## Slide 18
Conclusions and Outlook
  QCL & QEPAS: HONO detection, small volume
  LOD: 66 ppb (1 sec), 7 ppb (150 sec)
  DFG for analysis of surgical smoke (partially unknown composition)
  First quantitative in vivo analysis, identification of gases & anesthetic at ppm
ETH Zurich
Laser Spectroscopy and Sensing
  Mid-IR VECSEL for C1-C4 alkanes analysis
  Fast analysis at sub-ppm level
New laser developments:
ICLs, Diode-pumped Pb salt VECSELs, QCLs, frequency combs/supercontinuum
Detection schemes:
Cavity-enhanced techniques, ATR, QEPAS and cantilever PAS,integrated devices
New application areas
Field & POC measuremen, Lab-on-a-chip

## Slide 19
ETH Zurich
Laser Spectroscopy and Sensing
Acknowledgement
Julien Rey
Matthias Fill
Collaboration and Funding
Michele
Gianella
Ferdinand
Felder
Weidong Chen
Dr. Dieter Hahnloser

## Slide 20
Thank you for your attention
Notes: 20