## Slide 1: Intercomparison of nitrogen monoxide and nitrogen dioxide measurements in the atmosphere simulation chamber SAPHIR during the MetNO2 campaign
05th May 2020 I Robert Wegener and the MetNO2 SAPHIR intercomparison team
Franz Rohrer, Benjamin Winter, Zhaofeng Tan, Ulrich Bundke, Patrick Weber, Norbert Houben, Doreen Niether, Vaishali Vardhan, Sergej Wedel, Ralf Tillmann, Anna Novelli, Annika Kuß, Georg Stange, Robert Holla, Dagmar Kubistin, Karri Saarnio, Nicolas Sobanski, Claus Nordstroem, Jesper Nygaard, Tom Jensen, Naomi Farren, Simone Andersen, Will Drysdale, Sumin Kim, Sepyo Lee, Maitane Iturrate-Garcia, Sivan van Aswegen, David R. Worton
Contact: r.wegener@fz-juelich.de
Session AS5.11 Tuesday, 05 May 2020, 08:30-10:15

## Slide 2: Motivation
The MetNO2 Project
developing traceable static and dynamic reference standards
developing analytical methods to quantify the main impurities
The goal of the EMPIR Project MetNO2 within EMPIR is to
deliver a highly accurate SI traceable measurement of NO2.
NO2 is one of the key pollutants worldwide. Monitoring long-term trends requires accurate and traceable standards.
Methods and standards were compared at the SAPHIR chamber in Jülich
This will be achieved by
http://empir.npl.co.uk/metno2/

## Slide 3: Intercomparison SETUP
Instruments
Chemiluminescence Detectors (CLD) Photolytical converters Molybdenum converters |  | Univ of York (UK), Univ of Aarhus (DK), German Weather Service (GER), National Institute of Meteorological Sciences (South Korea) FZ Jülich (GER)
Cavity attenuated phase shift spectroscopy (CAPS) |  | Finnish Meteorological Institute (FI), German Weather Service (GER), Univ of York (UK), FZ Jülich (GER)
Quantum Cascade Laser (QCL) |  | Eidgenössische Materialprüfungs- und Forschungsanstalt (SUI) FZ Jülich (GER)
Iterative cavity-enhanced DOAS (ICAD) |  | FZ Jülich (GER)
Fontijn, A. 1971Journal of the Air Pollution Control Association, 21: 428-&.., 1970)
Kebabian etg al., 2008. , Environmental Science & Technology, 42: 6040-45.
Tuzson et al., 2013. Atmospheric Measurement Techniques, 6: 927-36.
Horbanski, et al., 2019. Atmospheric Measurement Techniques, 12: 3365-81.

## Slide 4: SETUP
Instruments
VS
SAPHIR
JULIAC
Instruments were set up in containers under the SAPHIR chamber and connected with a ½‘ PFA line.
Ambient air was provides from the tip of the JULIAC tower at 50 m height above a forest.
Photochemical experiments were conducted in the SAPHIR environmental chamber with a size of 270m3.

## Slide 5: SETUP
Instruments
VS
pump
NOx
PTRMS
Guest
Calibration
CAL
SAPHIR
JULIAC
Instruments
Instruments
MFC
Instruments
Sample air was provided via a common sample line either:
from a calibration unit
from SAPHIR
from the JULIAC tower

## Slide 6: SETUP
Intercomparison setup
Preliminary data were discussed daily.
Instruments were calibrated by the participants prior or after the campaign of two weeks.
During campaign, daily calibrations were performed via the common sampling line.
Calibration included a blank phase, an NO phase and an NO2 phase both under dry and humid conditions.
blank
Dry
NO
NO2
blank
Humid
NO
NO2
blank
Dry
NO
NO2
blank
Humid
NO
NO2
blank
Instruments were compared to a reference instrument calibrated with a traceable gas standard.

## Slide 7: RESULTS
Instrument sensitivities for NO
During the campaign instruments measured from the same standard with a mixing ratio of 14.8 ppb NO.
The QCL and some CLDs underestimated the concentration of the standard.
Time series of NO for the daily calibration
Average NO concentration measured during daily calibration
The mixing ratio of the standard is expressed with the dashed line
Many instruments showed changes in the NO values.

## Slide 8: RESULTS
Instrument sensitivities NO2
As for NO, also measured NO2 mixing ratios changed over the 14 days.
ICAD, QCLs and one CAPS agreed well with the reference instrument.
Time series of the daily calibration for NO2
Average NO2 concentration measured during daily calibration

## Slide 9: RESULTS
Instrument zero
While blank values in the NO channel of the CLDs were zero, some CLD showed a significant offset in the NO2 channel
Average NO2 offset during daily calibration
Average NO offset during daily calibration

## Slide 10: RESULTS
Standard comparisons
The standards of the participants were measured and compared to the reference instrument.
Measured and expected values agreed within ± 3 %.
Discrepancies of some instruments to the reference instrument cannot be explained with differences in the standards.
Recovery rate of the NO standards with respect to the value of the reference instrument.
Error bars show the uncertainties of the standards

## Slide 11: RESULTS
Standard comparisions
The provision of accurate NO2 standards is the main goal of MetNO2 project.
NO2 was provided from gas cylinders, produced by gas phase titration (GPT) from NO or permeation.
Two NO2 standards which were not provided from the MetNO2 project, differed by more than 10 %.
The standards from NPL and METAS developed within the project agreed well.
Recovery rate of the NO2 standards. Error bars show the uncertainties of the standards

## Slide 12: RESULTS
Influence of humidity
Water is an effective quenching substance in CLDs and reduces the sensitivity.
NO2 and NO data were scaled to the reference values at dry conditions to investigate the effect of humidity.
Data scaled to common standard
Dry
Humid
blank
NO
NO2
Blank
NO
NO2

## Slide 13: RESULTS
Influence of humidity
where [NO] is the measured and [NO]H2Ocorr the corrected concentration.
Recovery loss of NO standards due to water in percentage loss per absolute humidity
The value determined for the CLD is close to the value of 4.3 proposed by Parrish et al. (1991) and for suggested for the sensitivity correction in CLDs.
Parrish et al., 1991., Journal of Geophysical Research-Atmospheres, 96: 9309-20.
The influence of water on the sensitivity can be described with the a-value which is defined as:

## Slide 14: RESULTS
Influence of humidity
As the NO sensitivity, also the NO2 sensitivity in the CLDs dropped by 3-4% per water mixing ratio in %.
In contrast, the NO2 signal in one CAPS instrument was increased by the additional water in the cavity.
No effect was observed in the ICAD.
Recovery loss of NO2 standards due to water in percentage loss per absolute humidity

## Slide 15: RESULTS
Interferences from alkenes, nitrates and nitrites
Alkenes, nitrates and nitrites were added to the joint sampling line.
No response in the NO channels of the CLDs was observed after adding alkenes and nitrates.
Two CLDs showed interferences in the NO channel after adding isobutylnitrite.
NO time series when adding potential interfering compounds
The nitrite was used to simulate HONO interferences.

## Slide 16: RESULTS
Interferences from alkenes, nitrates and nitrites
The CLD with the molybdenum converter and, remarkably, a CLD with a photolytical converter measured NO2 after adding nitrates.
All CLDs showed interferences with isobutylnitrite.
The CAPS, the QCL and the ICAD did not show interferences.
NO2 time series when adding potential interfering compounds

## Slide 17: RESULTS
Ambient measurements
An Intercomparison of ambient NO2 measurements was performed by sampling from the JULIAC tower. NO2 values differed due to the scaling factors of the instruments. The differences could be reduced using the scaling factors from the joint calibrations.
Time series of ambient NO2 measurements.
Time series of ambient NO2 measurements scaled to joint calibrations and corrected for humidity effects.

## Slide 18: RESULTS
Ambient measurements
The ICAD showed an excellent correlation with the reference instrument with a slope of 1.
The slopes of the correlation plots of other instruments was mainly determined by the calibration and the humidity effect.

## Slide 19: RESULTS
When comparing with the reference instrument the slope is between 0.7 and 1.05.
For all but one instrument R2 is larger than 0.95.
Ambient measurements
After scaling to the joint reference, the slope is between 0.92 and 1.07.
Slope of all instruments with respect to the reference instrument. All instrumenents scaled to common standard and humidity corrected.

## Slide 20: SUMMARY
Sensitivity:
When measuring the calibration standards instruments deviated between -30% and +10%.
Humidity:
CLDs showed loss of sensitivity under humid conditions, but this effect could be corrected.
Reference standards:
Calibration worked well with the NO2 standards prepared within the MetNO2 project, but also with all the NO gas standards.
Interferences:
Not only molybdeneum but also one photolytical converter showed interferences with nitrates; all CLDs had interferences with isobutylnitrite and therefore potentially with HONO.
Ambient Measurements:
All instrument types showed good correlations with the reference instrument. Deviations can mainly be explained with the individual calibration procedures.
This intercomparison shows that the accuracy of NO and NO2 can greatly be improved by improving the calibration procedures. Calibration Gas mixtures with NO and CO2 should be used to track the dilution with precise CO2 measurements.

## Slide 21: THANK you
Forschungszentrum Jülich:
Franz Rohrer, Benjamin Winter, Zhaofeng Tan, Ulrich Bundke, Patrick Weber, Norbert Houben, Doreen Niether, Vaishali Vardhan, Sergej Wedel, Ralf Tillmann, Anna Novelli
German Weather Service (DWD)
Annika Kuß, Georg Stange, Robert Holla, Dagmar Kubistin
Eidgenössische Materialprüfungs- und Forschungsanstalt (EMPA)
Nicolas Sobanski
National Physical Laboratory NPL
Sivan van Aswegen, David R. Worton
Finnish Meteorological Institute
Karri Saarnio
University of Aarhus
Claus Nordstroem, Jesper Nygaard, Tom Jensen
University of York
Naomi Farren, Simone Andersen, Will Drysdale
National Institute of Meteorological Sciences (South Korea)
Sumin Kim, Sepyo Lee
Federal Institute of Metrology METAS
Maitane Iturrate-Garcia